## UKAZI S STRANI PROCESORSKE
Command from MT310 to MERILNA PCB, MT-300-TFA, HV ADAPTER
> | Transmitter | Receiver | : | ID | : | Function | : | Command | : | Parameter 1 | : | Parameter 2 | : | CRC flag | : | CRC | : | #013 | #010
char | char | char | char | string | char | string | char | string | char | string | char | string | char | char | char | string | char | char | char
Transmitter, Receiver
procesorska MT-310, P
merilna MT-310, M
adapter MT-300-TFA, A
HV adapter, H
ID
ID pošlje procesorska (zaenkrat dvomestna cifra), ostali v odgovoru isto vrnejo nazaj
Function
single funckija merilne, adapterja oz. HV adapterja, tabelo stringov funkcij je potrebno definirati
Command | glavni ukaz, kaj naj merilna oz. Adapter naredi
INIT | uporabljeno za inicializacijo funkcije, ob tem se lahko pošlje tudi parametre v Parameter 1 in Parameter 2, funkcije lahko ob tem klicu štartajo recimo V-metre za zunanjo napetost itd...
PARAMETER | pošiljanje parametrov, ob tem se ne izvede nič, samo shranijo se parametri
START | start meritve, za to morajo obstajati pogoji, za katere skrbi Procesor P
STOP | stop merive, če mora meritev ustaviti uporabnik, v primeru da se meritev ustavi avtomatsko oz. Na timer, tega ne rabimo
EXIT | izhod iz meritve, to procesor P pošlje, ko se recimo s preklopnikom izbere druga funkcija na preklopniku
Parameter 1
string Parameter 1 je sestavljen sledeče, PAR1_NAME¦PAR1_VALUE,PAR2_NAME¦PAR2_VALUE,itd...
PAR_NAME string, je ime parametra
PAR_VALUE string, je vrednost parametra, string za tekst oz. Število
Parameter 2
rezervirano za prihodnost, zaenkrat se ne pošilja nič
CRC flag
0 | ni potrebno izračunati CRC (checksum vseh znakov do vključno zadnjega dvopičja pred CRC stringom), vseeno se pošlje CRC string 000, sprejemnik to ignorira
1 | je potrebno izračunati CRC (checksum vseh znakov do vključno zadnjega dvopičja pred CRC stringom), po[lje se CRC
PRIMER
>PM:01:RCD:START:FUNC¦t,TYPE¦AC,CHAR¦GEN,MULT¦12x,CURRENT¦10,UC¦50,POL¦POS:00:1:023:#013#010 | >PM:01:RCD:START:FUNC¦t,TYPE¦AC,CHAR¦GEN,MULT¦12x,CURRENT¦10,UC¦50,POL¦POS:00:1:023:#013#010
prosesorska pošlje ukaz merilni
funkcija je RCD
ukaz je START
parametri so:
FUNC | t
TYPE | AC
CHAR | GEN
MULT | 12x
CURRENT | 10
UC | 50
POL | POS

## ODGOVORI
Answer from MERILNA PCB, MT-300-TFA, HV ADAPTER to MT-310 procesorska
> | Transmitter | Receiver | : | ID | : | Function | : | Command | : | Parameter 1 | : | Parameter 2 | : | CRC flag | : | CRC | : | #013 | #010
char | char | char | char | string | char | string | char | string | char | string | char | string | char | char | char | string | char | char | char
Transmitter, Receiver
procesorska MT-310, P
merilna MT-310, M
adapter MT-300-TFA, A
HV adapter, H
ID
ID (zaenkrat dvomestna cifra), vrnemo isto, kar je poslala procesorska
Function
funkcija, tabelo je potrebno definirati, da je ista v vseh procesorjih, katerih se to tiče
Command | tip odgovora
ACK | acknowledged
Parameter 1
string Parameter 1 je sestavljen sledeče, PAR1_NAME¦PAR1_VALUE,PAR2_NAME¦PAR2_VALUE,itd...
PAR_NAME string, je ime parametra
PAR_VALUE string, je vrednost parametra, string za tekst oz. Število
Parameter 2
rezervirano za prihodnost, zaenkrat se ne pošilja nič
CRC flag
0 | ni potrebno izračunati CRC (checksum vseh znakov do vključno zadnjega dvopičja pred CRC stringom), vseeno se pošlje CRC string 000, sprejemnik to ignorira
1 | je potrebno izračunati CRC (checksum vseh znakov do vključno zadnjega dvopičja pred CRC stringom), pošlje se CRC
PRIMER
>MP:01:RCD:VALUE:ULN:voltage:1:CRC:#013#010
merilna odgovori procesorski
funkcija je RCD
odgovor je VALUE
ime vrednosti je ULN
vrednost je voltage (string kot float število, decimalna pika)
ACKNOWLEDGE
za vsak ukaz, ki ga pošlje procesorska, merilna in adapterji vrnejo acknowlede odgovor, ki je, primer:
>MP:01:ACK::::1:CRC:#013#010
za stringom ACK so 4 dvopičja kot v odgovoru, vendar so tu polja Answer, Name, Value prazna
spodnje ne velja več, ker je sedaj format odgovora enka formatu ukaza s strani procesorske in se ves message pošlje kot en odgovor
ČE JE POTREBNO KOT ODGOVOR POSLATI VEČ VREDNOSTI
kadar se zgodi, da kot odgovor pošiljamo več vrednosti, npr. 3 medfazne napetosti, potem je dogovor, da vse tri rezultate pošljemo kot posamezen odgovor
zaenkrat sprogramirajmo tako, da lahko med ukaze damo nek enostavno generalno nastavljiv delay recimo 1ms (Janez, poglej in sporoč, če rabiš ta delay), v FT3 je bil
primer:
>MP:01:RCD:VALUE:ULN:voltage:1:CRC:#013#010
>MP:01:RCD:VALUE:ULPE:voltage:1:CRC:#013#010
>MP:01:RCD:VALUE:UNPE:voltage:1:CRC:#013#010
če se je zgodil EVENT, in recimo merilna pošlje asinhroni EVENT in potem še vrednosti, bi to zgledalo tako: (Janez preveri, kako sva rekla, ali je EVENT pred rezultati, ali na koncu)
>MP:01:RCD:EVENT:READY_ON::1:CRC:#013#010
>MP:01:RCD:VALUE:ULN:voltage:1:CRC:#013#010
>MP:01:RCD:VALUE:ULPE:voltage:1:CRC:#013#010
>MP:01:RCD:VALUE:UNPE:voltage:1:CRC:#013#010

## FUNKCIJE
GENERAL
VISUAL
RPE
LOOP
UDROP
RCD
RISO
HV_ADAPT
URES
ILEAK
UP
FUNC
AUTO

## UKAZI
ACK
ERROR
RESULT
START
GET
EXIT
VALUE
EVENT
STOP
CALIBRATION
INIT
VERSION
FAST
AUTO_STARTED
RESULTS
FUNC
TYPE
POL
CURRENT
MEAS
MODE
RAMP
STEP
SHORT
START_KEY_DOWN
START_KEY_UP
SAVE_KEY_DOWN
SAVE_KEY_UP
NEXT_KEY_DOWN
NEXT_KEY_UP
REMOTE_DISCONNECTED
PARAMETERS
1
0
MT_300_TFA_CONN
HV_ADAPT_CONN

## RCD
Command MT310
>PM:ID:RCD:INIT::00:1:023:#013#010
>PM:ID:RCD:START:FUNC¦t,TYPE¦AC,CHAR¦GEN,MULT¦12x,CURRENT¦10,UC¦50,POL¦POS:00:1:023:#013#010
>PM:ID:RCD:STOP::00:1:023:#013#010
>PM:ID:RCD:PARAMETER:FUNC¦t,TYPE¦AC,CHAR¦GEN,MULT¦12x,CURRENT¦10,UC¦50,POL¦POS:00:1:023:#013#010
>PM:ID:RCD:EXIT::00:1:023:#013#010
FUNC | TYPE | IN | MULT | UC | POL
t | AC | 6 mA | 1x | 50 | POS
| AC (S) | 10mA | 12x |  | NEG
| A | 30mA
| A (S) | 100mA
| B | 300mA
| B (S) | 500mA
1000mA
CUSTOM (6mA - 1000mA)
RAMP | AC | 6 mA
| AC (S) | 10mA
| A | 30mA
| A (S) | 100mA
| B | 300mA
| B (S) | 500mA
1000mA
CUSTOM (6mA - 1000mA)
Uc | AC | 6 mA
| AC (S) | 10mA
| A | 30mA
| A (S) | 100mA
| B | 300mA
| B (S) | 500mA
1000mA
CUSTOM (6mA - 1000mA)
AUTO | AC | 6 mA
| AC (S) | 10mA
| A | 30mA
| A (S) | 100mA
| B | 300mA
| B (S) | 500mA
1000mA
CUSTOM (6mA - 1000mA)
Returned answer
>MP:ID:RCD:ACK::00:1:023:#013#010
>MP:ID:RCD:VALUE:VOLTAGES¦,ULN¦,ULPE¦,UNPE¦:00:1:023:#013#010
>MP:ID:RCD:RESULT:UC¦,RA¦:00:1:023:#013#010
>MP:ID:RCD:EVENT:BAD_FUSE:00:1:023:#013#010
>MP:ID:RCD:ERROR::00:1:023:#013#010
COMMAND | COMMAND VALUE NAMES
RESULT | CALIBRATION
VOLTAGES
ULN
ULPE
UNPE
FREQUENCY
TRIPOUT_TIME
NO_TRIP_TIME
TRIP_CURRENT
RAMP_CURRENT
UC
RA
READY_ON
READY_OFF
READY_REVERSED
START_FAILED
EVENT | BAD_FUSE
MESUREMENT_FAILED
HOT_RESISTOR
HOT_TRANSISTOR
UC_OVER_LIMIT
N_DISCONNECTED
CURRENT_ERROR
MAINS_PE
CHAR
MULT

## RPE
Command MT310
>PM:ID:RPE:INIT::00:1:023:#013#010
>PM:ID:RPE:START:FUNC¦2W,MODE¦MAN,CURRENT¦10:00:1:023:#013#010
>PM:ID:RPE:STOP::00:1:023:#013#010
>PM:ID:RPE:PARAMETER:FUNC¦2W,MODE¦MAN,CURRENT¦10:00:1:023:#013#010
>PM:ID:RPE:EXIT::00:1:023:#013#010
FUNC | MODE | IN
2W | MAN | 200mA
| AUTO | 10A
25A
4W | MAN | 200mA
| AUTO | 10A
25A
2W+CLAMP | MAN | 200mA
| AUTO | 10A
25A
4W+CLAMP | MAN | 200mA
| AUTO | 10A
25A
Returned answer
>MP:ID:RPE:ACK::00:1:023:#013#010
>MP:ID:RPE:VALUE:VOLTAGES¦220,ULN¦20,ULPE¦21,UNPE¦22:00:1:023:#013#010
>MP:ID:RPE:RESULT:UC¦,RA¦:00:1:023:#013#010
>MP:ID:RPE:EVENT:BAD_FUSE_F32:00:1:023:#013#010
>MP:ID:RPE:ERROR::00:1:023:#013#010
COMMAND | COMMAND VALUE NAMES
AUTO_STARTED | INIT_OK
INIT_NOTOK
RESULT | RESISTANCE
I_CLAMP
RPE_TOT
I_MEAS
FAST_OK
FAST_NOTOK
EVENT | EXTERNAL_VOLTAGE
EXTERNAL_SENSE_VOLTAGE
END_EXTERNAL_VOLTAGE
BAD_FUSE_F32
CONN
MAN
NORMAL
LONG

## RISO
Command MT310
>PM:ID:RISO:INIT::00:1:023:#013#010
>PM:ID:RISO:START:MODE¦STEP,VOLTAGE¦50:00:1:023:#013#010
>PM:ID:RISO:STOP::00:1:023:#013#010
>PM:ID:RISO:PARAMETER:MODE¦STEP,VOLTAGE¦50:00:1:023:#013#010
>PM:ID:RISO:EXIT::00:1:023:#013#010
MEAS | MODE | UN | VOLTAGE_START | VOLTAGE_STOP | DURATION
NONE | STEP | 50V | 50 | 500 | POL
100V
250V
500V
1000V
CUSTOM (50V - 1000V)
| RAMP | 50V
100V
250V
500V
1000V
CUSTOM (50V - 1000V)
Returned answer
>MP:ID:RISO:ACK::00:1:023:#013#010
>MP:ID:RISO:VALUE:VOLTAGES¦50,RESISTANCE¦100:00:1:023:#013#010
>MP:ID:RISO:RESULT:UC¦,RA¦:00:1:023:#013#010
>MP:ID:RISO:EVENT:BAD_FUSE_F32:00:1:023:#013#010
>MP:ID:RISO:ERROR::00:1:023:#013#010
COMMAND | COMMAND VALUE NAMES
INIT_OK
INIT_NOTOK
RESULT | RESISTANCE
VOLTAGE
FAST_OK
FAST_NOTOK
VOLTAGE_START
VOLTAGE_STOP
DURATION
RAMP_RESULT
EVENT | EXTERNAL_SENSE_VOLTAGE
END_EXTERNAL_VOLTAGE
BAD_FUSE_F32

## UDROP
Command MT310
>PM:ID:UDROP:INIT::00:1:023:#013#010
>PM:ID:UDROP:START:ACC¦STEP,CURRENT¦50:00:1:023:#013#010
>PM:ID:UDROP:STOP::00:1:023:#013#010
>PM:ID:UDROP:PARAMETER:ACC¦STEP,CURRENT¦50:00:1:023:#013#010
>PM:ID:UDROP:EXIT::00:1:023:#013#010
MODE | ACC | IN | VOLTAGE
AUTO | ZL/N | 2A | 230V
4A
6A
8A
40A
12A
16A
20A
25A
32A
40A
50A
63A
80A
100A
125A
160A
200A
250A
315A
400A
500A
630A
| ZL/N-H | 2A
4A
6A
8A
40A
12A
16A
20A
25A
32A
40A
50A
63A
80A
100A
125A
160A
200A
250A
315A
400A
500A
630A
Returned answer
>MP:ID:UDROP:ACK::00:1:023:#013#010
>MP:ID:UDROP:VALUE:VOLTAGES¦50,RESISTANCE¦100:00:1:023:#013#010
>MP:ID:UDROP:RESULT:IMPEDANCE¦2500,VOLTAGE¦25:00:1:023:#013#010
>MP:ID:UDROP:EVENT:BAD_FUSE_F32:00:1:023:#013#010
>MP:ID:UDROP:ERROR::00:1:023:#013#010
COMMAND | COMMAND VALUE NAMES
INIT_OK
INIT_NOTOK
RESULT | IMPEDANCE
VOLTAGE
FAST_OK
FAST_NOTOK
RAMP_RESULT
EVENT | EXTERNAL_SENSE_VOLTAGE
END_EXTERNAL_VOLTAGE
BAD_FUSE_F32
ACC

## LOOP
Command MT310
>PM:ID:LOOP:INIT::00:1:023:#013#010
>PM:ID:LOOP:START:FUNC¦RA,MODE¦L/N/PE,CURRENT¦100:00:1:023:#013#010
>PM:ID:LOOP:STOP::00:1:023:#013#010
>PM:ID:LOOP:PARAMETER:FUNC¦RA,MODE¦L/N/PE,CURRENT¦100:00:1:023:#013#010
>PM:ID:LOOP:EXIT::00:1:023:#013#010
MEAS | MODE | IN | IL
STD/ZLN
STD/ZLPE
STD/ZLL
H/ZLN
H/ZLPE
H/ZLL
RA | L/N/PE | 6mA
| L/PE/PE | 10mA
30mA
100mA
300mA
500mA
1000mA
CUSTOM (6mA - 1000mA)
ULOW |  |  | CUSTOM (0A - 3A)
Returned answer
>MP:ID:LOOP:ACK::00:1:023:#013#010
>MP:ID:LOOP:VALUE:VOLTAGES¦50,RESISTANCE¦100:00:1:023:#013#010
>MP:ID:LOOP:RESULT:IMPEDANCE¦2500,VOLTAGE¦25:00:1:023:#013#010
>MP:ID:LOOP:EVENT:BAD_FUSE_F32:00:1:023:#013#010
>MP:ID:LOOP:ERROR::00:1:023:#013#010
COMMAND | COMMAND VALUE NAMES
INIT_OK
INIT_NOTOK
VALUE | I_LOAD,
I_P_S_C
UC
I_DIFF_NOM
EVENT | EXTERNAL_SENSE_VOLTAGE
END_EXTERNAL_VOLTAGE
BAD_FUSE_F32
ACC
RESULT | IMPEDANCE
VOLTAGE

## URES
Command MT310
>PM:ID:URES:INIT::00:1:023:#013#010
>PM:ID:URES:START:MODE¦STEP,VOLTAGE¦50:00:1:023:#013#010
>PM:ID:URES:STOP::00:1:023:#013#010
>PM:ID:URES:PARAMETER:MODE¦STEP,VOLTAGE¦50:00:1:023:#013#010
>PM:ID:URES:GET:VOLTAGE_IN:00:1:023:#013#010
>PM:ID:URES:FAST:MEAS¦URES,MODE¦LINEAR,VOLTAGE_NOM¦230,LIMIT¦1.0:00:1:023:#013#010
>PM:ID:URES:EXIT::00:1:023:#013#010
MEAS | MODE | LIMIT | LIMIT U
URES | LINEAR | 1s
| NONLINEAR | 5s
TRES |  |  | 60V
CUSTOM (10V - 650V)
Returned answer
>MP:ID:URES:ACK::00:1:023:#013#010
>MP:ID:URES:VALUE:VOLTAGE_IN¦100:00:1:023:#013#010
>MP:ID:URES:RESULT:U_RES¦60:00:1:023:#013#010
>MP:ID:URES:ERROR::00:1:023:#013#010
>MP:ID:LOOP:EVENT:REPEAT:00:1:023:#013#010
COMMAND | COMMAND VALUE NAMES
INIT_OK
INIT_NOTOK
LINEAR
NON_LINEAR
VOLTAGE_NOM
VALUE | VOLTAGE_IN
RESULT | U_RES
T_RES
LIMIT
READY_ON
READY_OFF
EVENT | REPEAT
TRIGGER
lk

## ILEAK
Command MT310
>PM:ID:ILEAK:INIT::00:1:023:#013#010
>PM:ID:ILEAK:START:MODE¦STEP,VOLTAGE¦50:00:1:023:#013#010
>PM:ID:ILEAK:STOP::00:1:023:#013#010
>PM:ID:ILEAK:PARAMETER:MODE¦STEP,VOLTAGE¦50:00:1:023:#013#010
>PM:ID:ILEAK:EXIT::00:1:023:#013#010
MEAS
ILEAK CLAMP
ILOAD CLAMP
ITOUCH
ILEAK W-PE
ILEAK P-PE
Returned answer
>MP:ID:ILEAK:ACK::00:1:023:#013#010
>MP:ID:ILEAK:RESULT:I_LEAK_CLAMP¦60:00:1:023:#013#010
>MP:ID:ILEAK:ERROR::00:1:023:#013#010
COMMAND | COMMAND VALUE NAMES
INIT_OK
INIT_NOTOK
RESULT | I_LEAK_CLAMP
I_LOAD_CLAMP
I_TOUCH
I_LEAK_W_PE
I_LEAK_P_PE
THD
FREQ
CURRENT
TIMER

## UP
Command MT310
>PM:ID:UP:INIT::00:1:023:#013#010
>PM:ID:UP:START:FUNC¦ULN,CONN¦4W:00:1:023:#013#010
>PM:ID:UP:STOP::00:1:023:#013#010
>PM:ID:UP:PARAMETER:FUNC¦PELV:00:1:023:#013#010
>PM:ID:UP:EXIT::00:1:023:#013#010
MEAS | CONN
UL/N | 2W
3W
4W
POWER
3PSEQ
PELV
SELV
URMS WELD
UPEAK WELD
UCONTROL
Returned answer
>MP:ID:UP:ACK::00:1:023:#013#010
>MP:ID:UP:RESULT:VOLTAGE¦60:00:1:023:#013#010
>MP:ID:UP:ERROR::00:1:023:#013#010
COMMAND | COMMAND VALUE NAMES
{L"INIT_OK"},
{L"INIT_NOTOK"},
{L"INTERUPTIONS"},
{L"ULN"},
{L"DIPS"},
{L"POWER"},
{L"3PSEQ"},
RESULT | {L"VOLTAGE"},
{L"f"},
{L"CONN"},
{L"2W"},
{L"THD"},
{L"UL-PE"},
{L"3W"},
{L"THD_L1-L2"},
{L"THD_L2-L3"},
{L"THD_L3-L1"},
{L"UL1-L2"},
{L"UL2-L3"},
{L"UL3-L1"},
{L"4W"},
{L"THD_L1-N"},
{L"THD_L2-N"},
{L"THD_L3-N"},
{L"UL1-N"},
{L"UL2-N"},
{L"UL3-N"},
{L"POWER"},
{L"IL"},
{L"Q"},
{L"COS_PHI"},
{L"U"},
{L"P"},
{L"PF"},
{L"PELV"},
{L"SELV"},
{L"IS_AC"},
{L"AC"},
{L"DC"},
{L"WELD_RMS"},
{L"WELD_PEAK"},
{L"CONTROL"},
{L"DIRECTION"},
{L"UZSC"},
{L"UNSC"},
{L"LEFT"},
{L"RIGHT"},

## HV-ADAPTER
Command MT310
>PM:ID:HV_ADAPT:INIT::00:1:023:#013#010
>PM:ID:HV_ADAPT:MT_READY:00:1:023:#013#010
>PM:ID:HV_ADAPT:STOP::00:1:023:#013#010
>PM:ID:HV_ADAPT:PARAMETERS:MODE:I_TRIPOUT,VOLTAGE¦2500,CURRENT¦1,CHAR¦APP,TIMER¦1,SAFETY1¦0,SAFETY_2¦0,WARN_LIGHR¦0:00:1:023:#013#010
>PM:HV_ADAPT:UP:EXIT::00:1:023:#013#010
MODE | VOLTAGE | CURRENT | CHAR | TIMER | SAFETY1 | SAFETY2
TRIPOUT | CUSTOM (250V - 5100V) | 1mA | IAPP | 1 | 0 | 0
|  |  | IREAL |  | 1 | 1
BURN
RAMP
PULSE
Returned answer
>MP:HV_ADAPT:UP:ACK::00:1:023:#013#010
>MP:ID:HV_ADAPT:RESULT:VOLTAGE¦60:00:1:023:#013#010
>MP:ID:HV_ADAPT:EVENT¦SELF_TEST_FAILED:00:1:023:#013#010
>MP:ID:HV_ADAPT:ERROR::00:1:023:#013#010
COMMAND | COMMAND VALUE NAMES
INIT_OK
INIT_NOTOK
RESULT | I_TEST
U_TEST
I_TRIPOUT
RAMP_50
BURN
CHAR
APP
REAL
FAST_OK
FAST_NOTOK
EVENT | HV_FUNCTION_LOCKED
SAFETY_1
SAFETY_2
WARN_LIGHT
PULSE
MT_READY
MT_NOT_READY
MEAS_STARTED
MEAS_STOPPED
TIMER
SELF_TEST_PASSED
SELF_TEST_FAILED
PRESS_START_AGAIN
PERFORM_SELFTEST
TIMER_STARTED

## MT-300-TFA
Command MT310 to MT300-TFA | Additional commands | Returned answer | Additional command | Command explained | Aditional parameter | Posible values
CABLE EXTENSIONS
MAIN FUNCTION CALL
[5 empty] | Phase system | "1P_SYS" | "3P_SYS"
>PA:ID:CORD:INIT:(1.)::1:CRC:#013#010 | 1. Cable type: "1_P" or "3_P" | >AP:ID:CORD:INITIATED:::1:CRC:#013#010 |  | init CORD test (if all OK returns INITIATED), runs mains test and return its result | Position of L conductor | "L" | "N"
|  | >AP:ID:CORD:INIT_FAILED:::1:CRC:#013#010 |  | init CORD not possible, initialisation failed
>PA:ID:CORD:DEINIT:::1:CRC:#013#010 |  | >AP:ID:CORD:DEINITIATED:::1:CRC:#013#010 |  | CORD deinitialized
>PA:ID:CORD:START_C_W:::1:CRC:#013#010 |  | >AP:ID:CORD:CW_STARTED:::1:CRC:#013#010 |  | start Correct wiring for CORD measurement | CW failures | L1_L2_CRIS | L1_L3_CRIS | L2_L3_CRIS | L1_N_CRIS | L1_PE_CRIS | L2_N_CRIS | L2_PE_CRIS | L3_N_CRIS | L3_PE_CRIS | N_PE_CRIS | L1_L2_SHOR | L1_L3_SHOR | L2_L3_SHOR | L1_N_SHOR | L1_PE_SHOR | L2_N_SHOR | L2_PE_SHOR | L3_N_SHOR | L3_PE_SHOR | N_PE_SHOR | MULTIFLT | CROS_N_SHOR | CROS_N_OPEN | SHOR_N_OPEN | MULTI_SHOR | MULTI_CROS | MULTI_OPEN | FAIL | PASS | L1_OPEN | L2_OPEN | L3_OPEN | N_OPEN | PE_OPEN
|  | >AP:ID:CORD:CW_RESULT:(1),(2)::1:CRC:#013#010 | 1. Pass 2. CW failures | Transmitt Cord Correct wiring results. When return message is recieved, measurement has finished. | Mains status | "OK" | "PE_DISC" | "N_DISC" | "WRONG_CON"
>PA:ID:CORD:START_CONT:::1:CRC:#013#010 |  | >AP:ID:CORD:CONT_STARTED:::1:CRC:#013#010 |  | start Continuity CORD measurement | Riso one phase to PE fault | "L1-PE_F" | "L2-PE_F" | "L3-PE_F" | "N-PE_F"
|  | >AP:ID:CORD:CONT:(1),(2)¦(3)::1:CRC:#013#010 | 1. Pass, 2. Measured wire, 3. Value in ohms %.3f | Transmitt Cord Continuity results. When return message is recieved, measurement has finished. | Riso phase to phase fault | "L1-N_F" | "L2-N_F" | "L3-N_F" | "L1-L3_F" | "L1-L2_F" | "L2-L3_F"
>PA:ID:CORD:START_ALL-PE:::1:CRC:#013#010 |  | >AP:ID:CORD:ALL-PE_STARTED:::1:CRC:#013#010 |  | start All Phases to PE CORD measurement | Measured wire | "L1" | "L2" | "L3" | "N" | "PE"
|  | >AP:ID:CORD:RISO_ALL-PE_RESULT:(1)::1:CRC:#013#010 | 1. Pass | Transmitt Cord All Phases to PE result. When return message is recieved, measurement has finished. | Pass | "PASS" | "FAIL"
>PA:ID:CORD:START_ONE-PE:::1:CRC:#013#010 |  | >AP:ID:ONE-PE_STARTED:::1:CRC:#013#010 |  | start Phases to PE CORD measurement | URES Phase voltage to measure | "L1-PE" | "L2-PE" | "L3-PE" | "L1-N" | "L2-N" | "L3-N"
|  | >AP:ID:CORD:RISO_ONE-PE_RESULT:(1),(2)::1:CRC:#013#010 | 1. Pass, 2. Riso one phase to PE fault-s | Transmitt Cord Riso Phases to PE result (up to 4 fault conditions). When return message is recieved, measurement has finished. | Device | "MT-310" | "MT-300-TFA" | "PAT"
>PA:ID:CORD:START_PH_PH:::1:CRC:#013#010 |  | >AP:ID:CORD:PH-PH_STARTED:::1:CRC:#013#010 |  | start Phases to Phase CORD measurement | Connection denied | "ALREADY_CONNECTED" | "DIFF_DEVICE_CONNECTED"
|  | >AP:ID:CORD:RISO_PH-PH_RESULT:(1),(2)::1:CRC:#013#010 | 1. Pass, 2. Riso phase to phase fault-s | Transmitt Cord Riso Phases to Phase result (up to 6 fault conditions). When return message is recieved, measurement has finished. | Cont/single | "SINGLE" | "CONT"
>PA:ID:CORD:SET:CABLE_TYPE¦(1.)::1:CRC:#013#010 | 1. Cable type: "1_P" or "3_P" | >AP:ID:CORD:SET_COMPLETED:::1:CRC:#013#010 |  | sets cable phase number (mandatory only if changed from INIT function)
>PA:ID:CORD:STOP:::1:CRC:#013#010 |  | >AP:ID:CORD:STOPPED:::1:CRC:#013#010 |  | stop and reinit CORD meaurement
INTERMEDIATE COMMANDS (MT-300-TFA asks)
>PA:ID:CORD:RPE_L_STARTED:::1:CRC: |  | >AP:ID:CORD:RPE_L_START:::1:CRC:#013#010 |  | Start RPE low current measurement
>PA:ID:CORD:RPE_H_STARTED:::1:CRC: |  | >AP:ID:CORD:RPE_H_START:::1:CRC:#013#010 |  | Start RPE high current measurement
>PA:ID:CORD:RPE_RES:RES¦(1.)::1:CRC: | 1. RPE in ohms %.2f | >AP:ID:CORD:RPE_RES_GET::(1):1:CRC:#013#010 | 1. Measured connection et.: "L1_L2" | Request RPE resistance. Parameter (1) just for info and simulation.
>PA:ID:CORD:RPE_STOPPED:::1:CRC: |  | >AP:ID:CORD:RPE_STOP:::1:CRC:#013#010 |  | Stop RPE measurement
>PA:ID:CORD:RISO_STARTED:::1:CRC: |  | >AP:ID:CORD:RISO_START:::1:CRC:#013#010 |  | Start RPE measurement
>PA:ID:CORD:RISO_RES:RES¦(1.)::1:CRC: | 1. RISO in Mohms %.3f | >AP:ID:CORD:RISO_RES_GET::(1):1:CRC:#013#010 | 1. Measured connection et.: "L1_L2" | Request RPE resistance. Parameter (1) just for info and simulation.
>PA:ID:CORD:RISO_STOPPED:::1:CRC: |  | >AP:ID:CORD:RISO_STOP:::1:CRC:#013#010 |  | Stop RISO measurement
|  | >AP:03:STATUS:MAINS:PHASE_S¦(1.),L_POS¦(2.),MAINS¦(3.)::1:CRC:#013#010 | 1. Phase system, 2. Position of L conductor, 3. Mains status | Mains power status
MACHINES
MAIN FUNCTION CALL
>PA:ID:MACH:INIT:(1.)::1:CRC:#013#010 | 1. Cable type: "1_P" or "3_P" | >AP:ID:MACH:INITIATED:::1:CRC:#013#010 |  | init MACH test, runs mains test and return its result
|  | >AP:ID:MACH:INIT_FAILED:::1:CRC:#013#010 |  | init MACH not possible, initialisation failed
>PA:ID:MACH:START_RPE:::1:CRC:#013#010 |  | >AP:ID:MACH:RPE_STARTED:::1:CRC:#013#010 |  | start RPE 30A MACH measurement, when return command is transmitted everithing is set
>PA:ID:MACH:STOP_RPE:::1:CRC:#013#010 |  | >AP:ID:MACH:RPE_STOPPED:::1:CRC:#013#010 |  | stop RPE 30A MACH measurement
>PA:ID:MACH:START_ALL-PE:::1:CRC:#013#010 |  | >AP:ID:MACH:ALL-PE_STARTED:::1:CRC:#013#010 |  | start All Phases to PE MACH measurement
|  | >AP:ID:MACH:RISO_ALL-PE_RESULT:(1)::1:CRC:#013#010 | 1. Pass | Transmitt MACH All Phases to PE result. When return message is recieved, measurement has finished.
>PA:ID:MACH:START_ONE-PE:::1:CRC:#013#010 |  | >AP:ID:MACH:ONE-PE_STARTED:::1:CRC:#013#010 |  | start Phases to PE MACH measurement. If ALL-PE is not completed jet, it is started implicitly.
|  | >AP:ID:MACH:RISO_ONE-PE_RESULT:(1),(2)::1:CRC:#013#010 | 1. Pass, 2. Riso one phase to PE fault-s | Transmitt MACH Riso Phases to PE result (up to 4 fault conditions). When return message is recieved, measurement has finished.
>PA:ID:MACH:START_URES:(1.):(2.):1:CRC:#013#010 | 1. URES Phase voltage to measure, 2. Parameters only for simulation | >AP:ID:MACH:URES_STARTED:::1:CRC:#013#010 |  | start dedicated URES measurement, when started is returned all is set
>PA:ID:MACH:URES_OPEN:::1:CRC:#013#010 |  | >AP:ID:MACH:URES_OPENED:::1:CRC:#013#010 |  | open input contactors at peak input voltage. Return when the contactors are opened (for info).
>PA:ID:MACH:URES_FINISHED:::1:CRC:#013#010 |  | >AP:ID:MACH:URES_STOPPED:::1:CRC:#013#010 |  | when URES measurement is finsihed in master device this string has to be returned
>PA:ID:MACH:SET:CABLE_TYPE¦(1.)::1:CRC:#013#010 | 1. Cable type: "1_P" or "3_P" | >AP:ID:MACH:SET_COMPLETED:::1:CRC:#013#010 |  | sets cable phase number (mandatory only if changed from INIT function)
>PA:ID:MACH:STOP:::1:CRC:#013#010 |  | >AP:ID:MACH:STOPPED:::1:CRC:#013#010 |  | stop and reinit MACH meaurement
>AP::::::0:000:#013#010
WELDING
MAIN FUNCTION CALL
>PA:ID:WELD:INIT:(1.)::1:CRC:#013#010 | 1. Cable type: "1_P" or "3_P" | >AP:ID:WELD:INITIATED:::1:CRC:#013#010 |  | init WELD test, runs mains test and return its result
|  | >AP:ID:WELD:INIT_FAILED:::1:CRC:#013#010 |  | init WELD not possible, initialisation failed
>PA:ID:WELD:START_RPE:::1:CRC:#013#010 |  | >AP:ID:WELD:RPE_STARTED:::1:CRC:#013#010 |  | start RPE 30A WELD measurement, when return command is transmitted everithing is set
>PA:ID:WELD:STOP_RPE:::1:CRC:#013#010 |  | >AP:ID:WELD:RPE_STOPPED:::1:CRC:#013#010 |  | stop RPE 30A WELD measurement
>PA:ID:WELD:START_ALL-PE:::1:CRC:#013#010 |  | >AP:ID:WELD:ALL-PE_STARTED:::1:CRC:#013#010 |  | start All Phases to PE WELD measurement
|  | >AP:ID:WELD:RISO_ALL-PE_RESULT:(1)::1:CRC:#013#010 | 1. Pass | Transmitt WELD All Phases to PE result. When return message is recieved, measurement has finished.
>PA:ID:WELD:START_ONE-PE:::1:CRC:#013#010 |  | >AP:ID:WELD:ONE-PE_STARTED:::1:CRC:#013#010 |  | start Phases to PE WELD measurement. If ALL-PE is not completed jet, it is started implicitly.
|  | >AP:ID:WELD:RISO_ONE-PE_RESULT:(1),(2)::1:CRC:#013#010 | 1. Pass, 2. Riso one phase to PE fault-s | Transmitt WELD Riso Phases to PE result (up to 4 fault conditions). When return message is recieved, measurement has finished.
>PA:ID:WELD:START_MAINS-WELD:(1.)::1:CRC:#013#010 | 1. Cont/single | >AP:ID:WELD:MAINS-WELD_STARTED:::1:CRC:#013#010 |  | start MAINS to welding WELD measurement. If measurement is set to continious started is returned when everithing is set and MT-300-TFA is not in command.
|  | >AP:ID:WELD:MAINS-WELD_RESULT:(1),(2)::1:CRC:#013#010 | 1. Pass | Transmitt WELD MAINS to welding result. When return message is recieved, measurement has finished.
>PA:ID:WELD:STOP_MAINS-WELD:::1:CRC:#013#010 |  | >AP:ID:WELD:MAINS-WELD_STOPPED:::1:CRC:#013#010 |  | stop MAINS to welding WELD measurement
>PA:ID:WELD:START_MAINS-CLASS2:(1.)::1:CRC:#013#010 | 1. Cont/single | >AP:ID:WELD:MAINS-CLASS2_STARTED:::1:CRC:#013#010 |  | start MAINS to CLASS2 WELD measurement. If measurement is set to continious started is returned when everithing is set and MT-300-TFA is not in command.
|  | >AP:ID:WELD:MAINS-CLASS2_RESULT:(1)::1:CRC:#013#010 | 1. Pass | Transmitt WELD MAINS to CLASS2 result. When return message is recieved, measurement has finished.
>PA:ID:WELD:STOP_MAINS-CLASS2:::1:CRC:#013#010 |  | >AP:ID:WELD:MAINS-CLASS2_STOPPED:::1:CRC:#013#010 |  | stop MAINS to CLASS2 WELD measurement.
>PA:ID:WELD:START_WELD-PE:(1.)::1:CRC:#013#010 | 1. Cont/single | >AP:ID:WELD:WELD-PE_STARTED:::1:CRC:#013#010 |  | start welding to PE WELD measurement. If measurement is set to continious started is returned when everithing is set and MT-300-TFA is not in command.
|  | >AP:ID:WELD:WELD-PE_RESULT:(1)::1:CRC:#013#010 | 1. Pass | Transmitt WELD welding to PE result. When return message is recieved, measurement has finished.
>PA:ID:WELD:STOP_WELD-PE:::1:CRC:#013#010 |  | >AP:ID:WELD:WELD-PE_STOPPED:::1:CRC:#013#010 |  | stop welding to PE WELD measurement
>PA:ID:WELD:START_UNL_RMS:::1:CRC:#013#010 |  | >AP:ID:WELD:UNL_RMS_STARTED:::1:CRC:#013#010 |  | start U-NO-LOAD RMS WELD measurement. When everithing is set started is returned
>PA:ID:WELD:STOP_UNL_RMS:::1:CRC:#013#010 |  | >AP:ID:WELD:UNL_RMS_STOPPED:::1:CRC:#013#010 |  | stop U-NO-LOAD RMS WELD measurement.
>PA:ID:WELD:START_UNL_PEAK:::1:CRC:#013#010 |  | >AP:ID:WELD:UNL_PEAK_STARTED:::1:CRC:#013#010 |  | start U-NO-LOAD PEAK WELD measurement. When everithing is set started is returned
>PA:ID:WELD:STOP_UNL_PEAK:::1:CRC:#013#010 |  | >AP:ID:WELD:UNL_PEAK_STOPPED:::1:CRC:#013#010 |  | stop U-NO-LOAD PEAK WELD measurement.
>PA:ID:WELD:START_ANALYZE:::1:CRC:#013#010 |  | >AP:ID:WELD:ANALYZE_STARTED:::1:CRC:#013#010 |  | start MAINS ANALYZW WELD measurement
>PA:ID:WELD:GET_VOLTAGE:::1:CRC:#013#010 |  | >AP:ID:WELD:VOLTAGE:ULN1¦(1.),ULN2¦(2.),ULN3¦(3.),UNPE¦(4.)::1:CRC:#013#010 | 1.-4. Voltage in %.2f | get measured voltage
>PA:ID:WELD:GET_CURRENT:::1:CRC:#013#010 |  | >AP:ID:WELD:CURRENT:IL1¦(1.),IL2¦(2.),IL3¦(3.)::1:CRC:#013#010 | 1.-3. Current in %.3f | get measured current
>PA:ID:WELD:GET_THD_C:::1:CRC:#013#010 |  | >AP:ID:WELD:THD_C:T_IL1¦(1.),T_IL2¦(2.),T_IL3¦(3.)::1:CRC:#013#010 | 1.-3. THD of current in %.2f | get current THD, updated every 2s
>PA:ID:WELD:GET_THD_V:::1:CRC:#013#010 |  | >AP:ID:WELD:THD_V:T_ULN1¦(1.),T_ULN2¦(2.),T-ULN3¦(3.)::1:CRC:#013#010 | 1.-3. THD of voltage in %.2f | get voltage THD, updated every 2s
>PA:ID:WELD:GET_POWER_R:::1:CRC:#013#010 |  | >AP:ID:WELDPOWER_R:PL1¦(1.),PL2¦(2.),PL3¦(3.),P3P¦(4.)::1:CRC:#013#010 | 1.-4. Power in %.1f | get computed real power
>PA:ID:WELD:GET_POWER_A:::1:CRC:#013#010 |  | >AP:ID:WELD:POWER_A:SL1¦(1.),SL2¦(2.),SL3¦(3.),S3P¦(4.)::1:CRC:#013#010 | 1.-4. Power in %.1f | get computed apparent power
>PA:ID:WELD:GET_PF:::1:CRC:#013#010 |  | >AP:ID:WELDPF:PFL1¦(1.),PFL2¦(2.),PFL3¦(3.),PF3P¦(4.)::1:CRC:#013#010 | 1.-4. Power factor in %.3f | get computed power factor
>PA:ID:WELD:STOP_ANALYZE:::1:CRC:#013#010 |  | >AP:ID:WELD:ANALYZE_STOPPED:::1:CRC:#013#010 |  | stop MAINS ANALYZW WELD measurement
>PA:ID:WELD:SET:CABLE_TYPE¦(1.)::1:CRC:#013#010 | 1. Cable type: "1_P" or "3_P" | >AP:ID:WELD:SET_COMPLETED:::1:CRC:#013#010 |  | sets cable phase number (mandatory only if changed from INIT function)
>PA:ID:WELD:STOP:::1:CRC:#013#010 |  | >AP:ID:WELD:STOPPED:::1:CRC:#013#010 |  | stop and reinit WELD meaurement
>AP::::::0:000:#013#010
CONNECTION
MAIN FUNCTION CALL
>PA:ID:COMMUN:CONNECT_REQUEST:(1.)::1:CRC:#013#010 | 1. Device | >PA:ID:COMMUN:CONNECTION_ESTABLED:(1.)::1:CRC:#013#010 | 1. Device | Connect request. Returns if all ok with Device (MT-300-TFA) , also returns mains status
|  | >PA:ID:COMMUN:CONNECTION_DENIED:(1.)::1:CRC:#013#010 | 1. Connection denied
>MP:01:RCD:EVENT:READY_ON::1:CRC:#013#010
>MP:01:RCD:EVENT:READY_OFF::1:CRC:#013#010
>MP:01:RCD:EVENT:READY_REVERSED::1:CRC:#013#010
>MP:01:RCD:EVENT:START_FAILED::1:CRC:#013#010
>MP:01:RCD:EVENT:BAD_FUSE::1:CRC:#013#010
>MP:01:RCD:EVENT:MESUREMENT_FAILED::1:CRC:#013#010